产品名称：; BSS称重传感器、BSS拉压传感器. 产品型号：. 产品类别：; 传感器系列 美国Transcell称重传感器. 产品简介：; BSS称重传感器、BSS拉压传感器，S型传感器；适用于材料试验机、配料秤、搅拌站、皮带秤和包装秤等.

产品描述

BSS称重传感器、BSS拉压传感器 产品简介INTRODUCTION： S型传感器；优质合金钢材质；安装简便、快速 ,稳定可靠；适用于材料试验机、配料秤、搅拌站、皮带秤和包装秤等； 量程Capacities: 100kg, 200kg, 250kg, 300kg, 500kg, 750kg, 1000kg, 1500kg, 2000kg, 2500kg, 5000kg, 10000kg 技术参数Specification: 额定输出; Rated Output; 3.0mV/V ±0.1%. 零点平衡; Zero Balance; ±1% of rated output. 蠕变; Creep after 30 minutes; ±0.03% of rated output. 非线性; Nonlinearity; ±0.03% of rated output. 滞后; Hysteresis; ±0.03% of rated output. 重复性; Repeatability; ±0.03% of rated output. 灵敏度温漂; Temp. effect on output; ≤0.002% of applied output/ºC. 零点温漂; Temp. effect on zero; ≤0.002% of rated output/ºC. 可用温度范围; Safe Temp. Range; -10ºC to + 70ºC. 温度补偿范围; Temp. Compensated; -10ºC to + 40ºC. 安全载荷; Safe Overload; 150%. 输入阻抗; Input Impedance; 387 ohm ± 20 ohm. 输出阻抗; Output Impedance; 350 ohm ± 5 ohm. 绝缘阻抗; Insulation Resistance; ≥5000 M ohm (50V DC). 推荐激励电压; Rated Excitation; 10V DC/AC. 最大激励电压; Maximum Excitation; 15V DC/AC. 导线长度; Cable Length; 6m. 导线颜色; Cable Color Code; Red(+E) Black(-E) Green(+S) White(-S).